## Slide 1: Extension of the operational capabilities of the Martin-Puplett Michelson Interferometer at W7-X
Hugo Isaac Cu Castillo

## Slide 2: Motivation
Electron temperature is one of the most important plasma parameters
Continuous monitoring necessary for:
  Electron heat transport, heat-pulse propagation, power balance studies, MHD phenomena, etc.
Why the Michelson?
  Broadband instrument -> probe several harmonics at the same time->X3 ECE available
  Radiometer has potential DC signal drifts
  Absolute calibration is possible
Calibration problem: huge front-end losses

## Slide 3: Outline
Electron Cyclotron Emission
Martin-Puplett Michelson Interferometer
Interferogram
Hot/cold-source calibration
Data acquisition
Characterization of Hot Source
Absolute calibration
Experiment worth looking at

## Slide 4: Electron Cyclotron Emission
Physical origin
Charged particle gyrates around B-field
Change of velocity = acceleration = radiation emission
modulated time retardation +dipole antenna +movement along
modulated time retardation +dipole antenna +movement along 𝑩 ⃗
Birth of harmonics
Charged particle gyrates around B-field
Change of velocity = acceleration = radiation emission
Birth of harmonics
O- and X- ECE modes appear
P=observer

## Slide 5: Electron Cyclotron Emission
Optical thickness: measure of how much the radiation is absorbed along a path
Cold resonance method:
Rayleigh-Jeans approximation of black body:
-Infinite optical thickness: -Connection intensity at a frequency to the electron temperature Optical thick modes at W7-X: O1, X2 Optically thin modes e.g. X3 also examinable (but cold resonance not valid)
Images: N. Chaudhary, “Investigation of optically grey electron cyclotron harmonics in Wendelstein 7-X”, 2020, PhD Thesis

## Slide 6: Martin-Puplett Michelson Interferometer
2 input – 2 output device
Beam splitter = wire grid polarizer
Rooftop mirrors rotate polarization 90°
x
Motor
Left: S. Schmuck et al, “Electron cyclotron emission spectra in X- and O-mode polarisation at JET: Martin-Puplett interferometer, absolute calibration, revised uncertainties, inboard/outboard temperature profile and wall properties,” Rev. Sci Instrum, vol. 87, 2016.
At W7-X:
mm
um
44 interferograms/s = samp rate: 22 ms
lid
Michelson
Front end-loses 20 dB
Right: N. Chaudhary, “Investigation of optically grey electron cyclotron harmonics in Wendelstein 7-X”, 2020, PhD Thesis

## Slide 7: Ideal interferogram
1. Monochromatic source: 2. Polychromatic source minus I:
4. Fourier Spectrum
Solve for I or C to find:
->
-> absolute calibration
… but one of them must be known before If I(f) from plasma is to be known, calibration is needed But S/N is low for calibration
3. Ideal interferogram
: amplification: calibration factors

## Slide 8: Real Interferogram
Additional terms:
: phase error
: source-dependent background rad.
: intensity-independent terms
: Gaussian noise
Varying power coupling of moving mirror at detector e.g.:
x0
x16 mm
Beam fixed mirror
Beam moving mirror
Zero path difference (zpd) not sampled

## Slide 9: Real Interferogram
If plasma discharge : is unknown, negligeable
: power coupling scales as
->polynomial fit
: 1) Estimate zpd
2) Apply FFT on (symmetric window function centered on estimated zpd)*((x) )
3)

## Slide 10: Real Interferogram
If calibration: is unknown, is not negligeable
Hot/cold source technique
Hot source
@873 K
Cold source
@77 K
Switch
1st, 2nd, 3rd harmonics

## Slide 11: Data acquisition
Mirror position and direction is encoded in I (blue) and Q (orange) signals,
Current acquisition scheme:
  Use output of interface electronics (green) to clock sampling ( um)
  One reference marker (purple) for determining mirror direction and for coherent averaging
Ref marker

## Slide 12: What has been done till now:
Obtention of calib. factors C(f) locally
  and are known: Eccosorb as black body emitter for low frequencies
Characterization of Hot Source i.e.
  C(f) is known for the local diagnostic
Eccosorb @297K
Eccosorb @77K (LN2)
Mirror
Towards Michelson
Hot Source @873 K

## Slide 13: What has been done till now:
Obtention of calib. factors locally
  I(f) is known: Eccosorb as black body emitter for low frequencies
Characterization of Hot Source
  C(f) is known for the local diagnostic
Water absorption lines
Eccosorb no longer black body
Insensitivity range for detector

## Slide 14: What must be done: Absolute calibration
Experiment:
  Save interferograms of Hot Source for 20 min
  Move blade (algorithm takes care of it)
  Save interferograms of Cold Source for 20 min
  Repeat for time required (hours/days)
  Repeat steps 1-4 for longer sections/more components
Controllable moving blade
Cold Source
@293 K?
@77 K?
Hot Source
@873
Do this locally and at TH
Towards Michelson
Data analysis:
  Filter electromagnetic noise
  Average coherently using data of ref. Marker
  Obtain FFT
  Work on statistics (std. dev)
Uncertainty decreases as N= number of interferograms
If front-end losses 20 dB = 100
-> factor 20 000 nr. of interferograms will be needed to achieve same S/N locally

## Slide 15: What is worth examining:
Radiation moves in waveguides in modes
And in the Michelson, only 10 Modes can reach detector (free space propagation)*
What if we only work with only fundamental mode?
Higher modes have anyways greater losses
Advantages:
  Commercial notch filter useable
  Stellartor: 1 Mode=better directivity of antenna
Possible compensation of signal loss: quasioptical sections of transmission line
reduction of mitre bends
*Assesment of Walter Kasparek
Free space propagation, p= mode number
Notes: Include measurements

## Slide 16: What is worth examining:
Radiation moves in waveguides in modes
And in the Michelson, only 10 Modes can reach detector (free space propagation)*
What if we only work with only fundamental mode?
Higher modes have anyways greater losses
Advantages:
  Commercial notch filter useable
  Stellartor: 1 Mode=better directivity of antenna
Possible compensation of signal loss: quasioptical sections of transmission line
reduction of mitre bends
*Assesment of Walter Kasparek
Notes: Include measurements

## Slide 17: Conclusions
Increase of S/N of calibration is now possible because:
  Longer integration times possible because of automatization
  Quality checks of obtained data
Improvement of instrument response
  Improvement of position of polarizer of spare port
Data augmentation (to investigate)

## Slide 18: Phase correction in detail

## Slide 19: Phase correction (0): Why necessary?*
Not sampling exactly the zpd introduces a phase error
If the interferogram would be sampled in :
But sampling is finite; Apodization function A(x) exists:
*More details: Porter, Tanner (1983): „Correction of phase errors in Fourier Spectroscopy“, International Journal of Infrared and Millimeter Waves, Vol. 4, No. 2

## Slide 20: Phase correction (0): Why necessary?
Convolution theorem: FFT of a product equals the convolution of the individual FFTs:
If is not symmetric, is complex:
Product results in mixing of real and imaginary parts (but it shouldn‘t!):
Solution: shift theorem: A(x) is translated by the right amount, resulting in:
Compensation possible

## Slide 21: Phase correction (1)
Information of zero path difference necessary:
Nr. of points for FFT = -> zero padding
m=index of maximum peak
<-zpd can be estimated from a parabola at max. 3 points

## Slide 22: Phase correction (2)
Window functions:
For phase correction For FFT analysis
(centereted around zpd)

## Slide 23: Phase correction (3) and FFT
Dispersion term at coarse grid
Finer grid:
Map via linear interpolation
FFT to find spectrum: